## Instructions
NFPA 99-2012 Risk Assessment Tool
Instructions for Using the ASHE NFPA 99 Risk Assessment Tool
Prior to implementing this risk assessment tool, the following steps should be taken:
1. Establish a multidisciplinary team with knowledge of the facility's space use, patient care services, clinical practices, and other areas as appropriate.
2. Familiarize the team with the risk category definitions found in chapters 4 (Fundamentals) and 12 (Emergency Management) of NFPA 99-2012: Health Care Facilities Code. These definitions are included in the category legends on each worksheet; mouse over the "Category Legends" box to see them.
3. Familiarize the team with the ways in which system and equipment operability can affect patient safety.
This risk assessment tool contains three worksheets (Systems, Equipment, and Emergency Management) as indicated on the worksheet tabs below.
Systems Worksheet - This worksheet is used to record the level of risk determined for the listed systems in a given area (room or spaces within a room) of the facility being evaluated. Indicate the risk level with an NFPA 99 risk category number (see the Category Legend for details).
Room Name: Enter the unique identification information for the room being evaluated (i.e., room name or number).
Room Number: Enter the room number, if applicable.
Space: Enter the unique identification information for the space in a room that is being evaluated (e.g., the charting area in a recovery area).
Chapter 5: Enter the risk category for the various components of the medical gas and vacuum systems in the room or space being evaluated.
Chapter 6: Enter the risk category for the electrical system in the room or space being evaluated.
Chapter 7: Enter the risk category for the various components of the IT and communications systems in the room or space being evaluated.
Chapter 8: Enter the risk category for the various components of the plumbing systems in the room or space being evaluated.
Chapter 9: Enter the risk category for the various components of the HVAC systems in the room or space being evaluated.
Chapter 10: Indicated on this worksheet for information only - to be assessed on the Equipment worksheet.
Chapter 12: Indicated on this worksheet for information only - to be assessed on the Emergency Management worksheet.
Note: Categories assigned in the chapter columns listed above are based on categories as outlined in Chapter 4 of NFPA 99-2012.
Equipment Worksheet - This worksheet is used to record the level of risk determined for all patient care equipment in the facility. Indicate the risk level with an NFPA 99 risk category number (see the Category Legend for details).
Equipment: Enter the name of the piece of equipment being assessed.
Equipment Tag #: Enter the unique identifying number for the piece of equipment being assessed. This information is optional, but should be considered when assessing non-movable equipment.
Category: Enter the risk category assigned to the equipment based on the categories outlined in Chapter 4 of NFPA 99-2012.
Notes: The comment area is provided for additional information about the reasons the risk category was assigned to the piece of equipment.
Note: NFPA 99-2012 (Section 1.3.1.1) defines the equipment covered by the code as “appliances and equipment used in patient care rooms of health care facilities.” This includes all equipment that may be used for patient care, such as defibrillators, ultrasound equipment, ventilators, weight scales, thermometers, and so on. Gas equipment includes items such as the oxygen regulator, non-rebreather masks, nebulizers, and nasal cannulas.
Emergency Management Worksheet - This tool is used to record the building category from NFPA 99 Table 12.3 (Application Matrix) assigned to each building.
Building: Enter the name or identifying information for the building.
Category: Enter the risk category assigned to each building based on the categories outlined in Table 12.3 (Application Matrix) in Chapter 12 of NFPA 99-2012.
Notes: This risk assessment tool has been developed to help health care facility staff comply with the risk-based, patient-focused approach required by NFPA 99: Health Care Facilities Code beginning with the 2012 edition. Rather than using the former occupancy-based approach, NFPA 99 now has the same requirements for a procedure no matter where it takes place, focusing on risks to patients and caregivers and on patient outcomes.
This completed risk assessment should be used to determine the steps needed to respond to the identified risks as outlined in NFPA 99. It should be kept as a record of the decisions made and updated annually.
v1.0

## Systems
Systems Risk Assessment Tool
Category Legend |  |  | Chapter 5 |  |  |  | Chapter 6 | Chapter 7 |  |  |  | Chapter 8 | [7 empty] | Chapter 9 |  |  | Chapters 10 and 11 |  |  | Chapter 12
Room Name | Room # | Space | Oxygen | Medical Air | Vacuum | WAGD | Electrical Systems | Data | Phone | Nurse Call | Cable TV | Potable Water | Non-Potable Water | Water Heating | Water Conditioning | Non-Medical Compressed Air | Black Waste Water | Gray Waste Water | Clear Waste Water | Heating | Ventilation | Air-Conditioning |  | Equipment | (See Equipment tab.) |  | Emergency Management | (See Emergency Management tab.)
Note on A3: Indicate the name of the room in this column.
Note on C3: When assessing the risk for multiple spaces in a room, use this column to identify each individual space.

## Equipment
Electrical and Gas Equipment Assessment Tool
Category Legend
Equipment | Equipment Tag # | Category | Notes
Patient moniters | Pre and post op areas | Electrical

## Emergency Management
Emergency Management Tool
| 1 | Those inpatient facilities that remain operable to provide advanced life support services to injured responders and disaster victims. These facilities manage the existing inpatient load as well as plan for the influx of additional patients as a result of an emergency.
| 2 | Those inpatient or outpatient facilities that augment the critical mission. These facilities manage the existing inpatient or outpatient loads but do not plan to receive additional patients as a result of an emergency.
Building | Category | Notes
| 2 | Patients will not be admitted during times of emergency, even if a waiver 1135 has been approved.
| 2 | No inpatient beds. Would discontinue elective cases during times of disaster.